# SUNA V2 Configuration Parameters
## Build Parameters
### Sensor Type
The Sensor Type is factory configured to either SUNA or ISUS.
An ISUS sensor uses a retroreflection probe for sampling the water, whereas the SUNA employs a gap in the sensor body to measure the absorption of the water sample.
### Sensor Version
The Sensor Version is factory configured to either FW, SW, Deep, APF, or some other custom version.
The FW (Fresh Water) version is calibrated for fresh water and is rated for perssures of up to xxx PSI.
The SW (Salt Water) version is calibrated for salt water and is rated for perssures of up to xxx PSI.
The Deep version is calibrated for salt water and is rated for perssures of up to xxx PSI.
The APF version is calibrated for salt water and is rated for perssures of up to xxx PSI.
### Serial Number
The Serial Number is factory configured.
### Lamp Shutter
The Lamp Shutter is factory configured to either Avaliable of Missing.
Normally, ISUS sensors utilize a lamp with a shutter, whereas SUNA sensors come without a shutter.
A lamp shutter allows for regular measurements of a dark sample, without switching off the lamp. it is of value when using the senor in long time continuous profiling applications. When using the sensor solely intermittendly for short durations, a shutter is of little additional value.
### Reference Detector
The Reference Detector is factory configured to either Avaliable of Missing.
Normally, ISUS sensors have a reference detector, whereas SUNA sensors do not.
The reference detector measures the bulk output of the sensor lamp. This measurement is used to ascertain that the lamp has been switched on. In a situation where the light path is blocked by some external obstruction (silt, bio-fouling), the reference intensity allows the operator to remotely determine if the sensor requires cleanup, or if the lamp is past its lifetime and has to be replaced.
### KBr Filter
The KBr Filter is factory configured to either Avaliable of Missing.
Normally, all sensors have a KBr filter.
The filter protects the optical components from UV damage by completely blocking light below 200 nm. As a side effect, the measurements below 200 nm can serve as a proxy for a dark sample.
### Super Capacitors
The Super Capacitors are factory configured to either Avaliable of Missing.
During startup, the capacitors are charged to provide brief internal power in the event that external power is suddenly cut off. Having some internal power allows the sensor to shut down into a safe and sane state.
The disadvantage of super capacitors is an increased total power consumption. Thus, in remote deployments, it may under certain conditions be advantageous to not use them.
### PCB Supervisor
The PCB Supervisor circuit is factory configured to either Avaliable of Missing.
Normally, all sensors are equipped with the power supervisor, which allows the sensor to enter power saving mode.
### USB Switch
The USB Switch is factory configured to either Avaliable of Missing.
Sensors can always communicate via RS-232. If USB is available and plugged in, communication switches to USB.
### Relay Board
The relay board circuit is factory configured to either Avaliable of Missing.
The relay board is only available in APF versions of the nitrate sensor,
### SDI-12 Board
The SDI-12 board circuit is factory configured to either Avaliable of Missing.
### Spectrometer Serial Number
The Spectrometer Serial Number is factory configured.
### Lamp Serial Number
The Lamp Serial Number is factory configured.
### Lamp Use Power
The power used by the lamp is determined during initial set-up of the sensor.
The lamp use power serves as a reference point, that allows the sensor to ascertain that the lamp is operating properly, and also provides the operator with an estimate of total power consumption.
### SDI-12 Address
The SDI-12 Address is configured internally by the system.
### Admin Password
The Administrator Password set to Nitrate.
Changing a few configuration parameters considered to be crucial requires the administrator password, to alert the operator that the change, if done inadvertendly, may negatively impact sensor performance.
## Output Control Parameters
### Baud Rate
The baud rate is one of 9600, 19200, 38400, 57600, or 115200.
A changed baud rate takes effect after the next power-up or reboot.
### Message Level
The message level is one of Error, Warn, Info, Debug, Trace.
Normally, the message level is set to Info.
Messages are directed to the output stream and are also saved in a message log file.
### Message File Size
The message file size is in the 0 to 65 MB range.
Setting the file size to zero turns off logging of messages to file.
This setting takes effect after the next reboot.
### Data File Size
The data file size is in the 1 to 64 MB range.
Normally, the data file size is set to 2 MB.
This value applies only if the data file type is set to continuous. Dayly and per-acquisition files will contain as much data as is generated.
### Output Frame Type / Logging Frame Type
The frame type is one of None, APF, MBARI, Full_ASCII, Full_Binary, Reduced_Binary, Concentration.
If set to None, no frame data will be written to serial output / data log file, respectively.
APF and MBARI are provided for the APEX Float version only.
### Output Dark Frame / Logging Dark Frame
Dark frames outout and logging is either Output or Suppress.
This configuration flag is provided if dark frames are not required or desired.
### Log File Type
The data log file type is one of Acquisition, Continuous, or Daily.
All data log files names have a single letter (A, C, or D) followed by a 7-digit number, followed by a 3-letter extension.
Acquisition based data files are started new whenever external power was cycled.
Continuous data log files are appended to until the Data File Size is reached. Then, the file number is incremented, and data are appended to the next file.
Daily data log files contain all data that are collected within a 24 hour period. The 7-digit number is made up of a 4-digit year and a 3-digit day-of-year (1 to 365 or 366 for leap years).
### DAC Minimum / Maximum Nitrate
TODO
### Dat Wavelength Low / High
APF version only. Set to 217 and 250.
These wavelength define the channels that are included in the APF frame.
## Operational
### Operation Mode
The operation mode is Continuous, FixedTime, Periodic, Polled, APF, or SDI12.
In Continuous mode the sensor starts to acquire data as soon as initialization is complete and countdown has expired. Data acquisition proceeds, depending on the Operation Control setting, either in a sample based (1 dark sample, then Light Samples, Dark Samples, Light Samples, ...) or time based (1 dark sample, then Light Duration, Dark Duration, Light Duration) infinite cycle.
In Fixedtime mode, the sensor behaves as in Continuous mode, only it terminates after a maximum of Fixed Time Duration seconds.
In Periodic mode, the sensor acquires data in regular periods, and collects, depending on the Operation Control setting, either a fixed number of sampls (Periodic Samples) or for a fixed time (Periodic Duration).
In Polled mode, the sensor stays in low power sleep, to acquire data only when receiving commands from an external controller.
In APF mode, the sensor stays in low power sleep, to acquire data only when receiving commands from an external controller.
SDI-12 mode brings the sensor in conformace with the SDI-12 interface specification.
### Operation Control
The operation control is Duration or Samples based.
Operation control applies to continuous, fixed time, and periodic mode. Either of these operating modes is further controlled by additional parameters, and this parameter determines which parameters apply.
### External Device
The external device is None, Wiper, or Pump.
### ExDev Pre Run Time
The external device pre-acquisition run time is measured in seconds, and can take any positive value up to 120.
### ExDev On At Acquire
The external device is powered On or Off during the acquisition.
For a pump, this flag must be On.
### Watchdog
The watchdog can be On of Off.
The watchdog is factory configured to be On.
### Countdown
The coundown is measured in units of seconds.
It is used in Continuous and FixedTime operation modes.
### Fixed Time Duration
The fixed time duration is measured in units of seconds, and can take any positive number up to and including 1000000.
### Periodic Interval
The periodic interval is restricted to values that divide the day into integer parts: 1m, 2m, 5m, 6m, 10m, 15m, 20m, 30m, 1h, 2h, 3h, 4h, 6h, 8h, 12h, 24h.
### Periodic Offset
The periodic offset is measured in seconds.
Whereas the periodic interval establishes a grid of acquisition times, the offset locates the grid relative to the start of the day (hour 0).
Note: There is a side effect when an external device needs to run prior to data acquisition.
### Periodic Duration
The periodic duration is measured in seconds.
This parameter is used when Operation Control is set to Duration.
### Periodic Samples
The periodic offset is measured in number of frames.
This parameter is used when Operation Control is set to Samples.
### Polled Timeout
The polled timeout is measured in seconds.
It determines for how long the firmware will wait for a command upon wakeup before returning to low power state.
### APF auto turn off timeout
The APF auto turn off timeout is measured in hours.
This setting applies to APF mode only.
### Reference Min At LampOn
Only used if Reference Detector is Available.
The reference min at lamp-on is factory configured and measured in counts.
When switching on the lamp, the reference detector must register at least the specified number of counts. Otherwise, the firmware assumes a lamp failure, tries a few more times, and switches itself off if it cannot see sufficient light.
### Clean Burn Cycle Duration
This setting applies to APF mode only.
## Acquisition
### Spectrometer Integration Period
TODO...
### Dark Averages
TODO...
### Light Averages
TODO...
### Reference Samples
TODO...
### PreAcquisition Scans
Obsolete...
### Dark Samples
TODO...
### Light Samples
TODO...
### Dark Duration
TODO...
### Light Duration
TODO...
### AcqTime of Last Sample
APEX Float Only. (Last == Most Recent)
### Number Of Samples
TODO...
## Processing Parameters
### Temperature Compensation
The temperature compensation flag is On or Off.
Real-time processing temperature compensation only works for saltwater calibrated sensors that are equipped to receive temperature and salinity input. The flag will be ignored if the sensor is not able to perform this task.
### Salinity Fitting
The salinity fitting flag is On or Off.
Salinity fitting can only be switched off in saltwater calibrated sensors that are equipped to receive temperature and salinity input. The flag will be ignored if the sensor is not able to perform this task.
### Bromide Tracing
The bromide tracing flag is On or Off.
Bromide tracing only works in freshwater environs. The sensor must either be a freshwater sensor, or if it is a seawater caklibrated sensor, the number of concentrations to fit must be set to 1. The bromide estimate will be reported in the output frame in a dedicated field. When attempting to trace bromide in freshwater, it may be advantageous to adjust the fitting interval, depending on the expected bromide concentrations.
### Concentrations to Fit
The number of concentrations to fit is 1, 2, or 3.
Freshwater calibrated sensors only have 1 concentration to fit; saltwater calibrated sensors can be made to act like freshwater sensor by setting the concentrations to fit to 1. Normally, saltwater calibrated sensors operate with 3 fitted concentrations.
### Baseline Order
The baseline order is fixed to 1.
Historically, different baseline orders were avaliable in Satlantic'sNitrate Sensors. However, there is no longer any need to change the baseline order.
### Dark Correction Method
The dark correction method is one of SpecAverage, SWKBrAverage, or FormulaBased.
The purpose of dark correction is to subtract the temperature dependent dark baseline from the measurement. When using SpecAverage, a dark spectrum is measured by either closing the shutter (of present) or switching off the lamp. Using the SWKBrAverage works if seawater of bromide cause extinction below 200 nm, and the measurement in that wavelength range is used as a proxy for the dark baseline. An experimental FormulaBased approach calculates the dark baseline from spectrometer integration time and sensor temperature.
### Dark Coefficients
The four dark coefficients characterize the dark baseline behaviour of the spectrometer. They are obtained in a calibration procedure.
### Absorbance Threshold
The absorbance threshold is a value between 0.01 and 10.0. It is normally set to 1.2.
Whenever the absorbance of a channel exceeds the specified absorbance threshold, that channel is excluded from processing. Setting the threshold to the maximum value of 10.0 will guarantee that all channels will be included in processing.
### Adaptive Integration Time
The absorbance threshold is a value between 0.01 and 10.0. It is normally set to 1.2.
Whenever the absorbance of a channel exceeds the specified absorbance threshold, that channel is excluded from processing. Setting the threshold to the maximum value of 10.0 will guarantee that all channels will be included in processing.
### Fitting Interval
The fitting interval is normally from 217 to 240 nm.
Changing the fitting interval should be done with caution; an unsuitable fitting interval can generate invalid results.