SUNA V2 and SUNA/ISUS APF
Firmware Installation and Initial Configuration Setup
### Atmel JTEG Board: Driver and Supporting Software Installation
\\Stiletto\Designs\Firmware\Atmel\Tools\AVR32_GNU_Toolchain\avr32-gnu-toolchain-2.4.2-setup.exe
### Installation of PCB Supervisor
Not on APF sensors. To be determined [Scott Feener]
### Installation of Bootloader:
Connect JTEG to main nitrate board.
Execute \\Stiletto\Designs\Firmware\Atmel\NitrateMotherboard\Bootloader\Vx.y.z\boot.bat to install the bootloader.
Disconnect JTEG after this step.
Connecting via RS232, 57600 baud, 8 bit, no parity, no handshake to PC.
Power cycle, and be ready to input 'b' key within 3 seconds.
Expect to see on serial port
>>>>>
Enter 'b' to break into bootloader
Expect to see
Satlantic Serial Bootloadera
Version x.y.z
SATBLDR>
Type “?” for help.
Now ready to install firmware.
### Installation of Firmware:
Latest firmware is located at \\Stiletto\Designs\Firmware\Atmel\NitrateMotherboard\Firmware
In bootloader, give “w” command, then send firmware “sfw” file via XMODEM.
After completion of upload,
give “a” command (instructs bootloader to execure application upon startup),
then “s” command or power cycle to start application firmware.

### Setup of Sensor:
When the firmware executed for the first time, it recognizes that nothing has been initialized or configured, and enters setup mode.Setup mode requests a list of inputs, and on the way runs a substancial set of tests.
The firmware prints “Entering system setup”,
and then requests operator confirmation “Hit key to enter.”
The data flash and eMMC card may not yet be formatted. If they are not formatted, the operator should format them. After formatting both, the sensor will reboot, and re-enter setup. The second time, instead of a request to format, disk sizes are shown.
Then follows a list of configuration values:
Clock Time	YYYY/MM/DD hh:mm:ss
Admin password:	APF:	DUNLIN
Other:	Nitrate
Factory password:	SAT+NO3
Sensor Type:	ISUS | SUNA
Sensor Version:	V2
Serial Number	between 1 and 9999
Super-Caps	Available | Missing	(can use 'a' or 'm' for short)
PCB Supervisor	Available | Missing
Relay Board	Available | Missing
SDI-12 Board	Available | Missing
USB Switch	Available | Missing
Scan for 1-wire sensors	May disconnect one to see which is which.
Fiberllight Serial Number	the serial number
Lamp Shutter	Available | Missing
Lamp Time Used	0 [seconds]
KBr Filter	Available | Missing
Reference Detector	Available | Missing
If Reference Detector	Firmware scans for intensity of reference detector
Set Reference Limit = 1/4 of those counts
Spectrometer Serial Number	Enter the serial number without leading zero
Upload ZCOEF file	via XMODEM
Loop that permits adjusting spectrometer gain.
N	to skip
Y	to cycle
Scan for best spectometer integration period; takes a while.
N	to skip
Y	to scan
Set integration period	Value from scan
Firmware measures dark, light spectrum and reports average, noise, min, max.
Dark average should be below 1000.
Dark noise must be below 10.
Also get lamp power use (should be 4500..5500 mW).
Build dummy cal file	Takes a while.
Test low-power sleep	N	to skip
Y	to test
Test 1: wait 30 seconds for from RTC wake-up
Test 2: enter a key to trigger wakeup from telemetry
Set burnin hours	0:	No burnin at reboot / power cycle
>0:	Will burnin for given hours after reboot / power cycle.
Setup prints configured values, and requests setting out-of-range values to default.
Y	confirm to set values to default
Setup enters command shell.
TODO (Burkhard): Specify which values to adjust.
Power cycle or Reboot to start burn-in.
### Burn-In (2012-05-07):
Burn-in is done
without the lamp-side housing installed,
with a fan providing air-flow over the transformers of the lamp assembly
Burn-in duration is 50 hours of lamp time.
Burn-in cycle is a long on interval, interrupted by a brief off interval,
3300 seconds on (=1 hour minus 300 seconds)
300 seconds off
The firmware will automatically execute the burn-in cycle when the factory parameter burnhours is set to a value >0.
Burn-in data are evaluated for
dark noise
light signal-over-noise ratio
nitrate noise (within the last 30-60 seconds of each on-interval)
nitrate drift (monitored at the end of each on-interval)
Data are collected internally, but output can also be logged to file.
After completion, all files from EVL folder are offloaded to \\Stiletto\SUNA\SUNA####\Burnin.
Burnin evaluation TODO (Burkhard). Note: A burn-in summary is generated on the sensor.